High efficiency - low cost.

Customised test loads and easy handling.

N3A

N3D G3

Rockwell

acc. to EN ISO 6508, ASTM E-18

HRA

HRB

HRC

HRD

HRE

HRF

HRG

HRH

HRK

HRL

HRM

HRP

HRR

HRS

HRV

HR 15-N

HR 30-N

HR 45-N

HR 15-T

HR 30-T

HR 45-T

HR 15-W

HR 30-W

HR 45-W

HR 15-X

HR 30-X

HR 45-X

HR 15-Y

HR 30-Y

HR 45-Y

Plastic test

acc. to EN ISO 2039

49.03 N

132.9 N

357.9 N

961 N

HVD, HBD

Depth measurement

HVD 30

HVD 100

1/30

2.5/31.5

2.5/62.5

2.5/187.5

5/62.5

2

N3 product brochure

长  春  恒  越  电  子  科  技  有  限  公  司

www.jlhykj.com 0431-86796799   www.jlhydz.net 0431-86796758

Easy operation.

Load application by means of 180° lever

(pre- and main load step).

1 Application of pre-load

2 Application of main load

3 Application of 2nd pre-load

4 Back to initial position

Digital and convenient.

The three free selectable test methods and conversions stored as stan-

dard in the tester for maximum flexibility. The reading is shown on the

backlit LCD screen after the test. Selection of 5 language versions

(GE, EN, IT, FR, SP). Machine power supply via mains adaptor or via USB

interface (laptop).

Tolerance limits can also be stored in the tester. Readings designated

with LOW, HIGH or OK accordingly. This is a big advantage, particularly

for serial part testing. The system immediately detects the parts to be

removed. Furthermore, it is also possible to test cylindrical and spheri-

cal work pieces. When the diameter is entered the hardness value adapts

corresponding to the convex surface.

Data export and data management.

All hardness values can be imported into EMCO-TEST software ecos via

RS232 or USB interface and easily managed. It is possible to generate

voluminous and meaningful statistics to be printed out as an A4 report.

Data can also be transfered directly to an excel document via USB.

N3 product brochure

3

长  春  恒  越  电  子  科  技  有  限  公  司

www.jlhykj.com 0431-86796799   www.jlhydz.net 0431-86796758

Approved technology.

Changeable spring sleeve system for variable test loads.

Variable test loads.

exchangeable

spring sleeves

The N1A test unit with exchangeable spring sleeves is used for different

loads. (Spring sleeves with different loads).

Spring sleeve N2H

Spring sleeve N2C

Spring sleeve N2G

Test unit

N1A

Spring sleeve N2L

Spring sleeve N2E

Spring sleeve N2B

Spring sleeve N2F

Spring sleeve N2A

Spring sleeve N2D

Test load 15 kgf

Test load 30 kgf

Test load 31.25 kgf

Test load 45 kgf

Test load 60 kgf

Test load 62.5 kgf

Test load 100 kgf

Test load 150 kgf

Test load 187.5 kgf

Several test loads combined.

Four load adjustable test units available. Required test load selected via

lever.

Test unit N1U*

Test unit N1R

Test unit N1S

Test unit N1P

*For N3D G3 only

Selectable test loads 30/150/187.5 kgf

Selectable test loads 60/100/150 kgf

Selectable test loads 15/30/45 kgf

Selectable test loads 5/13.5/36.5/98 kgf

Dial gauge.

N3A hardness tester measurement readings via analogue dial gauge.

Three coloured scales printed on the clock-face. Hardness value read

on corresponding scale according to chosen test method and test load

(accuracy 0.5).

N3 product brochure

4

长  春  恒  越  电  子  科  技  有  限  公  司

www.jlhykj.com 0431-86796799   www.jlhydz.net 0431-86796758

Accessories.

Complete accessory catalogue at www.emcotest.com

All N3 hardness testing machine accessories at www.emcotest.com: various penetrators, special nose cones, plane and V-anvils, spring sleeves,

load adjustable test units, spindle-bellow, data cable, test blocks and a good deal more.

Penetrators

Plane anvils and v-anvils

Technical data.

All basic data at a glance.

N3D G3

Measurement readings

Height adjustment:

Test height (N1A unit and spring sleeves):

Test height (with load adjustable test unit):

Throat depth:

Max. work piece weight:

Dimensions w x h x d:

Basic machine weight:

Data interface:

Power supply:

USB mains adaptor:

Max. power consumption:

Room temperature:

Digital

Hand wheel

285 mm

230 mm

170 mm

20 kg

205 x 830 x 470 mm

35 kg

1x RS232, 1x USB

110/230~1/N/PE

5 V DC

5W

5 - 40°

Hand wheel

285 mm

230 mm

170 mm

20 kg

N3A

Analogue via dial gauge

205 x 920 x 470 mm

35 kg

-

-

-

-

5 - 40°

N3 product brochure

5

长  春  恒  越  电  子  科  技  有  限  公  司

www.jlhykj.com 0431-86796799   www.jlhydz.net 0431-86796758